July 19, 2005

TO: 		Lillian Reichenthal

STEREO Instrument Manager

FROM: 	545/Thermal Engineering/J. Hawk

SUBJECT: 	SIT/SEPT – Thermal Vacuum and Thermal Balance Test Report

This report is to provide the results of the Suprathermal Ion Telecope (SIT), and the Solar Electron Proton Telescope (SEPT) thermal balance (TB) tests. Also included in this report are the results of the SIT instruments thermal vacuum (TV) tests. The tests were conducted in the Dynavac thermal vacuum chamber, building 7 at the NASA GSFC.

These Instrument level tests were conducted to fulfill the thermal vacuum requirement for the SEPT and SIT instruments.  The tests consisted of two thermal balance phases with a chamber break in between.  Subsequently, the SIT instruments underwent thermal vacuum cycling. SEPT had already completed the required instrument level TV cycling.  The Phase 1 thermal balance test was conducted from February 10 through February 17, 2005 and included test articles SEPT-E Ahead, SEPT-E Behind and SIT Ahead.  The Phase 2 thermal balance test was conducted from February 21 through March 9, 2005 and included test articles SEPT-NS Ahead, SEPT-NS Behind and SIT Behind.  Next, the AHEAD and BEHIND SIT instruments underwent thermal vacuum cycling. Shown in Figures 1 and 2 is the Phase 1 thermal balance test set-up. Figures 4 and 5 show the Phase 2 thermal balance test set-up. The SIT thermal vacuum test set-up is shown in Figure 6. During the SIT Ahead TB the instrument temperatures were significantly warmer than predicted (~20°C). After the Phase 1 TB chamber break, it was discovered that the thin film thermal coating for the radiators were applied with the low emissivity inconel side facing out instead of the high emissivity ITO silver Teflon side. Inconel’s low emissivity explained the significantly warmer than predicted temperatures. The thermal coating was removed and replaced correctly with the ITO silver Teflon side facing out and TB was repeated on SIT Ahead.

Thermal Balance Results

Data from the TB test was used to correlate the thermal model.  After minor changes to the thermal model thermal blanket effectiveness and isolator conductance couplings the temperature predictions matched test results within an average of 2ºC.

Flight predictions were made with the correlated model for the three cases:

- Cold Survival

- Cold Operational

- Hot Operational

The test setup was carried out according to the STEREO-GSFC-002.C test plan. The assemblies were mounted to the interface plate using the flight isolator feet. Flight-like MLI was installed in the flight configuration.  The radiator surfaces have the flight high emissivity ITO silver Teflon installed to dissipate heat. Ground straps between the test article and the interface plate were configured as flight-like as possible.  Flight cables were routed and secured to the interface plate as flight-like as possible.  The cables, which in flight run along the spacecraft deck, were coiled up and attached to the interface plate under the MLI blanket. GSE cables ran from the flight cables to the chamber feed-thru connecter and to external GSE.

Figure 1 SIT/SEPT Phase 1 TB Test Set-up

Figure 2 SIT/SEPT Phase 1 TB Test

Figure 3 SIT AHEAD TB Retest Set-up

Figure 4 SIT/SEPT Phase 2 TB Test Set-up

Figure 5 SIT/SEPT Phase 2 TB Test

TABLE 1: Phase 1 Comparison Test vs. Analysis Results (°C)

| Survival Cases |
| --- |
| SEPT AHEAD Ecliptic |
| TC # | Location | Measured | NODE # | Predicts | DELTA T |
| TC24 | Front of SEPTA on Bracket Ecliptic | -12.9 | ASEPT.93 | -14.0 | 1.1 |
| TC25 | Front of SEPTA Telescope Ecliptic | -13.0 | ASEPT.64 | -14.5 | 1.6 |
| TC26 | Front of SEPTA Telescope Ecliptic | -12.2 | ASEPT.21 | -15.3 | 3.1 |
| TC27 | Front of SEPTA Telescope Ecliptic | -12.3 | ASEPT.21 | -15.3 | 3.0 |
| TC28 | Front of SEPTA on Bracket Ecliptic | -12.3 | ASEPT.106 | -14.7 | 2.5 |
| TC29 | Front of SEPTA on Bracket Ecliptic | -13.0 | ASEPT.94 | -14.1 | 1.1 |
| TC30 | Front of SEPTA EBOX Ecliptic | -12.6 | ASEPT.35 | -14.6 | 2.0 |
| TC31 | Back of SEPTA on Bracket Ecliptic | -13.1 | ASEPT.105 | -14.6 | 1.5 |
| TC32 | Back of SEPTA Telescope Ecliptic | -13.2 | ASEPT.82 | -14.0 | 0.8 |
| TC33 | Back of SEPTA Telescope Ecliptic | -13.7 | ASEPT.28 | -13.0 | 0.6 |
| TC34 | Back of SEPTA Telescope Ecliptic | -13.9 | ASEPT.28 | -13.0 | 0.8 |
| TC35 | Back of SEPTA on Bracket Ecliptic | -13.8 | ASEPT.92 | -13.9 | 0.1 |
| TC36 | Back of SEPTA on Bracket Ecliptic | -12.6 | ASEPT.109 | -14.6 | 2.1 |
| TC37 | Back of SEPTA EBOX Ecliptic | -12.6 | ASEPT.40 | -14.6 | 1.9 |
| TC38 | Bottom of SEPTA Bracket Ecliptic | -12.0 | ASEPT.109 | -14.6 | 2.7 |
| TC39 | Bottom of SEPTA Bracket Ecliptic | -12.8 | ASEPT.94 | -14.1 | 1.4 |
| TC40 | Radiator Side of SEPTA Ecliptic | -13.6 | ASEPT.39 | -14.4 | 0.8 |
| TC106 | Bottom of SEPTA Ecliptic | -13.5 | ASEPT.36 | -14.4 | 0.9 |
| TC107 | Bottom of SEPTA Ecliptic | -12.4 | ASEPT.36 | -14.4 | 2.0 |
| SEPT BEHIND Ecliptic |
|  |  | Measured | NODE | Predicts | DELTA T |
| TC108 | Front of SEPTB on Bracket | -27.8 | BSEPT.93 | -27.2 | 0.6 |
| TC109 | Front of SEPTB Telescope | -27.6 | BSEPT.64 | -26.7 | 0.9 |
| TC110 | Front of SEPTB Telescope | -27.7 | BSEPT.21 | -26.4 | 1.3 |
| TC111 | Front of SEPTB Telescope | -27.6 | BSEPT.21 | -26.4 | 1.2 |
| TC112 | Front of SEPTB on Bracket | -27.1 | BSEPT.106 | -27.2 | 0.1 |
| TC113 | Front of SEPTB on Bracket | -27.9 | BSEPT.94 | -27.3 | 0.6 |
| TC114 | Front of SEPTB EBOX | -27.2 | BSEPT.35 | -27.4 | 0.3 |
| TC115 | Back of SEPTB on Bracket | -27.1 | BSEPT.105 | -27.3 | 0.2 |
| TC116 | Back of SEPTB Telescope | -27.2 | BSEPT.82 | -27.5 | 0.3 |
| TC117 | Back of SEPTB Telescope | -27.2 | BSEPT.28 | -27.6 | 0.4 |
| TC118 | Back of SEPTB Telescope | -27.3 | BSEPT.28 | -27.6 | 0.3 |
| TC119 | Back of SEPTB on Bracket | -27.7 | BSEPT.92 | -27.2 | 0.5 |
| TC120 | Back of SEPTB on Bracket | -26.9 | BSEPT.109 | -27.3 | 0.5 |
| TC121 | Back of SEPTB EBOX | -27.4 | BSEPT.40 | -27.5 | 0.1 |
| TC122 | Bottom of SEPTB Bracket | -26.9 | BSEPT.109 | -27.3 | 0.5 |
| TC123 | Bottom of SEPTB Bracket | -28.3 | BSEPT.94 | -27.3 | 1.0 |
| TC126 | Radiator Side of SEPTB EBOX | -28.8 | BSEPT.39 | -27.3 | 1.5 |
| TC124 | Bottom of SEPTB | -27.5 | BSEPT.36 | -27.1 | 0.4 |
| TC125 | Bottom of SEPTB | -26.7 | BSEPT.36 | -27.1 | 0.4 |
| Cold Operational Cases |
| SEPT AHEAD Ecliptic |
| TC # | Location | Measured | NODE # | Predicts | DELTA T |
| TC24 | Front of SEPTA on Bracket Ecliptic | -6.2 | ASEPT.93 | -9.1 | 2.9 |
| TC25 | Front of SEPTA Telescope Ecliptic | -6.2 | ASEPT.64 | -9.8 | 3.6 |
| TC26 | Front of SEPTA Telescope Ecliptic | -5.5 | ASEPT.21 | -10.7 | 5.2 |
| TC27 | Front of SEPTA Telescope Ecliptic | -5.6 | ASEPT.21 | -10.7 | 5.1 |
| TC28 | Front of SEPTA on Bracket Ecliptic | -5.4 | ASEPT.106 | -9.8 | 4.4 |
| TC29 | Front of SEPTA on Bracket Ecliptic | -6.3 | ASEPT.94 | -9.0 | 2.7 |
| TC30 | Front of SEPTA EBOX Ecliptic | -5.4 | ASEPT.35 | -8.8 | 3.4 |
| TC31 | Back of SEPTA on Bracket Ecliptic | -6.3 | ASEPT.105 | -9.7 | 3.4 |
| TC32 | Back of SEPTA Telescope Ecliptic | -6.5 | ASEPT.82 | -9.2 | 2.7 |
| TC33 | Back of SEPTA Telescope Ecliptic | -7.0 | ASEPT.28 | -8.4 | 1.4 |
| TC34 | Back of SEPTA Telescope Ecliptic | -7.2 | ASEPT.28 | -8.4 | 1.2 |
| TC35 | Back of SEPTA on Bracket Ecliptic | -7.1 | ASEPT.92 | -8.9 | 1.9 |
| TC36 | Back of SEPTA on Bracket Ecliptic | -5.6 | ASEPT.109 | -9.6 | 3.9 |
| TC37 | Back of SEPTA EBOX Ecliptic | -5.5 | ASEPT.40 | -8.8 | 3.3 |
| TC38 | Bottom of SEPTA Bracket Ecliptic | -5.1 | ASEPT.109 | -9.6 | 4.5 |
| TC39 | Bottom of SEPTA Bracket Ecliptic | -6.0 | ASEPT.94 | -9.0 | 3.0 |
| TC40 | Radiator Side of SEPTA Ecliptic | -6.9 | ASEPT.39 | -8.9 | 2.0 |
| TC106 | Bottom of SEPTA Ecliptic | -7.0 | ASEPT.36 | -9.1 | 2.1 |
| TC107 | Bottom of SEPTA Ecliptic | -5.6 | ASEPT.36 | -9.1 | 3.5 |
| SEPT BEHIND Ecliptic |
| TC108 | Front of SEPTB on Bracket | -20.2 | BSEPT.93 | -18.0 | 2.2 |
| TC109 | Front of SEPTB Telescope | -20.0 | BSEPT.64 | -17.7 | 2.3 |
| TC110 | Front of SEPTB Telescope | -20.1 | BSEPT.21 | -17.7 | 2.4 |
| TC111 | Front of SEPTB Telescope | -19.9 | BSEPT.21 | -17.7 | 2.2 |
| TC112 | Front of SEPTB on Bracket | -19.3 | BSEPT.106 | -18.1 | 1.2 |
| TC113 | Front of SEPTB on Bracket | -20.2 | BSEPT.94 | -18.0 | 2.2 |
| TC114 | Front of SEPTB EBOX | -19.1 | BSEPT.35 | -17.5 | 1.6 |
| TC115 | Back of SEPTB on Bracket | -19.3 | BSEPT.105 | -18.2 | 1.2 |
| TC116 | Back of SEPTB Telescope | -19.5 | BSEPT.82 | -18.6 | 0.9 |
| TC117 | Back of SEPTB Telescope | -19.6 | BSEPT.28 | -18.9 | 0.7 |
| TC118 | Back of SEPTB Telescope | -19.7 | BSEPT.28 | -18.9 | 0.8 |
| TC119 | Back of SEPTB on Bracket | -20.0 | BSEPT.92 | -18.1 | 1.9 |
| TC120 | Back of SEPTB on Bracket | -19.0 | BSEPT.109 | -18.1 | 1.0 |
| TC121 | Back of SEPTB EBOX | -19.3 | BSEPT.40 | -17.5 | 1.8 |
| TC122 | Bottom of SEPTB Bracket | -18.9 | BSEPT.109 | -18.1 | 0.8 |
| TC123 | Bottom of SEPTB Bracket | -20.5 | BSEPT.94 | -18.0 | 2.5 |
| TC126 | Radiator Side of SEPTB EBOX | -21.1 | BSEPT.39 | -17.6 | 3.5 |
| TC124 | Bottom of SEPTB | -19.9 | BSEPT.36 | -17.7 | 2.3 |
| TC125 | Bottom of SEPTB | -18.8 | BSEPT.36 | -17.7 | 1.1 |
| Hot Operational Cases |
| SEPT AHEAD Ecliptic |
|  |  | Measured | NODE # | Predicts | DELTA T |
| TC24 | Front of SEPTA on Bracket Ecliptic | 24.0 | ASEPT.93 | 20.8 | 3.2 |
| TC25 | Front of SEPTA Telescope Ecliptic | 24.1 | ASEPT.64 | 19.9 | 4.2 |
| TC26 | Front of SEPTA Telescope Ecliptic | 25.1 | ASEPT.21 | 18.6 | 6.5 |
| TC27 | Front of SEPTA Telescope Ecliptic | 25.2 | ASEPT.21 | 18.6 | 6.6 |
| TC28 | Front of SEPTA on Bracket Ecliptic | 25.1 | ASEPT.106 | 19.6 | 5.5 |
| TC29 | Front of SEPTA on Bracket Ecliptic | 23.7 | ASEPT.94 | 20.6 | 3.1 |
| TC30 | Front of SEPTA EBOX Ecliptic | 24.6 | ASEPT.35 | 20.0 | 4.6 |
| TC31 | Back of SEPTA on Bracket Ecliptic | 23.9 | ASEPT.105 | 19.8 | 4.1 |
| TC32 | Back of SEPTA Telescope Ecliptic | 23.6 | ASEPT.82 | 20.9 | 2.8 |
| TC33 | Back of SEPTA Telescope Ecliptic | 22.9 | ASEPT.28 | 22.4 | 0.5 |
| TC34 | Back of SEPTA Telescope Ecliptic | 22.6 | ASEPT.28 | 22.4 | 0.2 |
| TC35 | Back of SEPTA on Bracket Ecliptic | 22.7 | ASEPT.92 | 21.0 | 1.7 |
| TC36 | Back of SEPTA on Bracket Ecliptic | 24.5 | ASEPT.109 | 19.7 | 4.8 |
| TC37 | Back of SEPTA EBOX Ecliptic | 24.3 | ASEPT.40 | 20.0 | 4.3 |
| TC38 | Bottom of SEPTA Bracket Ecliptic | 25.1 | ASEPT.109 | 19.7 | 5.3 |
| TC39 | Bottom of SEPTA Bracket Ecliptic | 23.9 | ASEPT.94 | 20.6 | 3.3 |
| TC40 | Radiator Side of SEPTA Ecliptic | 22.5 | ASEPT.39 | 20.0 | 2.4 |
| TC106 | Bottom of SEPTA Ecliptic | 22.7 | ASEPT.36 | 19.6 | 3.1 |
| TC107 | Bottom of SEPTA Ecliptic | 24.6 | ASEPT.36 | 19.6 | 5.0 |
| SEPT BEHIND Ecliptic |
|  |  | Measured | NODE # | Predicts | DELTA T |
| TC108 | Front of SEPTB on Bracket | 2.0 | BSEPT.93 | 5.8 | 3.8 |
| TC109 | Front of SEPTB Telescope | 2.1 | BSEPT.64 | 6.6 | 4.5 |
| TC110 | Front of SEPTB Telescope | 2.1 | BSEPT.21 | 6.9 | 4.8 |
| TC111 | Front of SEPTB Telescope | 2.3 | BSEPT.21 | 6.9 | 4.6 |
| TC112 | Front of SEPTB on Bracket | 3.0 | BSEPT.106 | 5.7 | 2.8 |
| TC113 | Front of SEPTB on Bracket | 1.6 | BSEPT.94 | 5.6 | 4.0 |
| TC114 | Front of SEPTB EBOX | 2.8 | BSEPT.35 | 5.4 | 2.6 |
| TC115 | Back of SEPTB on Bracket | 3.0 | BSEPT.105 | 5.6 | 2.6 |
| TC116 | Back of SEPTB Telescope | 2.9 | BSEPT.82 | 5.2 | 2.3 |
| TC117 | Back of SEPTB Telescope | 2.9 | BSEPT.28 | 5.0 | 2.1 |
| TC118 | Back of SEPTB Telescope | 2.8 | BSEPT.28 | 5.0 | 2.3 |
| TC119 | Back of SEPTB on Bracket | 2.1 | BSEPT.92 | 5.7 | 3.6 |
| TC120 | Back of SEPTB on Bracket | 3.4 | BSEPT.109 | 5.5 | 2.0 |
| TC121 | Back of SEPTB EBOX | 2.7 | BSEPT.40 | 5.4 | 2.7 |
| TC122 | Bottom of SEPTB Bracket | 3.5 | BSEPT.109 | 5.5 | 2.0 |
| TC123 | Bottom of SEPTB Bracket | 1.3 | BSEPT.94 | 5.6 | 4.3 |
| TC126 | Radiator Side of SEPTB EBOX | 0.2 | BSEPT.39 | 5.4 | 5.2 |
| TC124 | Bottom of SEPTB | 1.8 | BSEPT.36 | 5.2 | 3.5 |
| TC125 | Bottom of SEPTB | 3.4 | BSEPT.36 | 5.2 | 1.8 |

TABLE 2: Phase 2 Comparison Test vs. Analysis Results (°C)

| Survival Cases |
| --- |
| SIT BEHIND |
| TC # | Location | Measured | NODE # | Predicts | DELTA T |
| TC1 | Front of SIT EBOX Behind | -29.8 | BSIT.444 | -27.0 | 2.9 |
| TC2 | Front of HVPS Behind | -30.0 | BPOWER.325 | -27.1 | 2.9 |
| TC3 | Right Side of SIT EBOX Behind | NA | BSIT.431 | -27.0 | NA |
| TC4 | Right Side of HVPS Behind | -30.1 | BPOWER.328 | -27.1 | 3.0 |
| TC5 | Right Side of SIT Telescope Behind | -30.7 | BSIT.435 | -27.9 | 2.7 |
| TC6 | Top of SIT Telescope Behind | -30.0 | BSIT.433 | -28.3 | 1.8 |
| TC7 | Top Front of SIT Telescope Behind | -31.8 | BSIT.425 | -29.2 | 2.6 |
| TC8 | Top of SIT Collimator Behind | -51.2 | BSIT.437 | -45.3 | 5.9 |
| TC9 | Right Side of SIT Collimator Behind | -49.5 | BSIT.419 | -45.0 | 4.5 |
| TC10 | Front of SIT Telescope Behind | -31.4 | BSIT.429 | -28.4 | 3.0 |
| TC11 | Top of SIT Telescope Behind | -30.5 | BSIT.433 | -28.3 | 2.3 |
| TC12 | Back of SIT Telescope Behind | -30.5 | BSIT.415 | -27.8 | 2.7 |
| TC13 | Left Side of SIT Ebox Behind | -29.4 | BSIT.431 | -27.0 | 2.4 |
| TC14 | Left Side of HVPS Behind | -29.8 | BPOWER.327 | -27.1 | 2.7 |
| TC15 | Left Side of SIT Telescope Behind | -31.5 | BSIT.413 | -28.0 | 3.6 |
| TC16 | Front of SIT Telescope Behind | -32.1 | BSIT.429 | -28.4 | 3.6 |
| TC17 | Back of SIT Telescope Behind | -30.8 | BSIT.415 | -27.8 | 3.1 |
| TC18 | Back of HVPS Behind | -30.1 | BPOWER.326 | -27.1 | 2.9 |
| TC19 | Back of SIT EBOX Behind | -29.5 | BSIT.409 | -27.0 | 2.6 |
| TC20 | Bottom of SIT EBOX Behind | -29.5 | BSIT.432 | -26.9 | 2.6 |
| TC21 | Bottom of SIT EBOX Behind | -29.6 | BSIT.432 | -26.9 | 2.7 |
| TC22 | Under SIT Collimator exterior Behind | -51.5 | BSIT.427 | -45.2 | 6.3 |
| TC23 | Right Side of SIT Collimator Behind | -50.6 | BSIT.419 | -45.0 | 5.5 |
| SEPT AHEAD NS |
| TC # | Location | Measured | NODE | Predicts | DELTA T |
| TC24 | Front of SEPTA on Bracket NS | -20.0 | SEPTNS.93 | -20.8 | 0.9 |
| TC25 | Front of SEPTA Telescope NS | -20.1 | SEPTNS.64 | -21.1 | 1.0 |
| TC26 | Front of SEPTA Telescope NS | -20.3 | SEPTNS.21 | -22.0 | 1.7 |
| TC27 | Front of SEPTA Telescope NS | -20.3 | SEPTNS.21 | -22.0 | 1.7 |
| TC28 | Front of SEPTA on Bracket NS | -19.1 | SEPTNS.106 | -21.0 | 1.9 |
| TC29 | Front of SEPTA on Bracket NS | -19.3 | SEPTNS.94 | -19.3 | 0.0 |
| TC30 | Front of SEPTA EBOX NS | -16.6 | SEPTNS.35 | -15.2 | 1.4 |
| TC31 | Back of SEPTA on Bracket NS | -19.7 | SEPTNS.105 | -20.9 | 1.2 |
| TC32 | Back of SEPTA Telescope NS | -19.7 | SEPTNS.82 | -21.1 | 1.4 |
| TC33 | Back of SEPTA Telescope NS | -20.7 | SEPTNS.28 | -21.9 | 1.3 |
| TC34 | Back of SEPTA Telescope NS | -20.7 | SEPTNS.28 | -21.9 | 1.2 |
| TC35 | Back of SEPTA on Bracket NS | -19.8 | SEPTNS.92 | -21.0 | 1.1 |
| TC36 | Back of SEPTA on Bracket NS | -18.5 | SEPTNS.109 | -19.3 | 0.8 |
| TC37 | Back of SEPTA EBOX NS | -16.4 | SEPTNS.40 | -15.2 | 1.2 |
| TC38 | Bottom of SEPTA Bracket NS | -16.9 | SEPTNS.109 | -19.3 | 2.4 |
| TC39 | Bottom of SEPTA Bracket NS | -18.0 | SEPTNS.94 | -19.3 | 1.3 |
| TC40 | Radiator Side of SEPTA NS | -16.9 | SEPTNS.39 | -15.2 | 1.7 |
| TC106 | Bottom of SEPTA NS | -16.0 | SEPTNS.36 | -13.5 | 2.5 |
| TC107 | Bottom of SEPTA NS | -14.6 | SEPTNS.36 | -13.5 | 1.1 |
| SEPT BEHIND NS |
| TC # | Location | Measured | NODE | Predicts | DELTA T |
| TC108 | Front of SEPTB on Bracket NS | -16.6 | SEPTNSB.93 | -21.1 | 4.5 |
| TC109 | Front of SEPTB Telescope NS | -17.0 | SEPTNSB.64 | -21.4 | 4.4 |
| TC110 | Front of SEPTB Telescope NS | -17.1 | SEPTNSB.21 | -22.2 | 5.1 |
| TC111 | Front of SEPTB Telescope NS | -16.9 | SEPTNSB.21 | -22.2 | 5.3 |
| TC112 | Front of SEPTB on Bracket NS | -15.9 | SEPTNSB.106 | -21.2 | 5.3 |
| TC113 | Front of SEPTB on Bracket NS | -15.1 | SEPTNSB.94 | -19.6 | 4.5 |
| TC114 | Front of SEPTB EBOX NS | -12.9 | SEPTNSB.35 | -15.4 | 2.5 |
| TC115 | Back of SEPTB on Bracket NS | -16.2 | SEPTNSB.105 | -21.1 | 4.9 |
| TC116 | Back of SEPTB Telescope NS | -16.7 | SEPTNSB.82 | -21.4 | 4.7 |
| TC117 | Back of SEPTB Telescope NS | -17.7 | SEPTNSB.28 | -22.3 | 4.6 |
| TC118 | Back of SEPTB Telescope NS | -17.8 | SEPTNSB.28 | -22.3 | 4.4 |
| TC119 | Back of SEPTB on Bracket NS | -16.1 | SEPTNSB.92 | -21.3 | 5.1 |
| TC120 | Back of SEPTB on Bracket NS | -17.1 | SEPTNSB.109 | -19.6 | 2.4 |
| TC121 | Back of SEPTB EBOX NS | -14.0 | SEPTNSB.40 | -15.4 | 1.5 |
| TC122 | Bottom of SEPTB Bracket NS | -15.4 | SEPTNSB.109 | -19.6 | 4.2 |
| TC123 | Bottom of SEPTB Bracket NS | -15.2 | SEPTNSB.94 | -19.6 | 4.4 |
| TC126 | Radiator Side of SEPTB EBOX NS | -13.7 | SEPTNSB.39 | -15.4 | 1.7 |
| TC124 | Bottom of SEPTB NS | -10.6 | SEPTNSB.36 | -13.8 | 3.2 |
| TC125 | Bottom of SEPTB NS | NA | SEPTNSB.36 | -13.8 | NA |
| Cold Operational Cases |
| SIT BEHIND |
| TC # | Location | Measured | NODE # | Predicts | DELTA T |
| TC1 | Front of SIT EBOX Behind | 0.7 | BSIT.444 | -1.7 | 2.4 |
| TC2 | Front of HVPS Behind | 0.3 | BPOWER.325 | -1.7 | 1.9 |
| TC3 | Right Side of SIT EBOX Behind | NA | BSIT.431 | -1.6 | NA |
| TC4 | Right Side of HVPS Behind | 0.1 | BPOWER.328 | -1.6 | 1.7 |
| TC5 | Right Side of SIT Telescope Behind | -9.6 | BSIT.435 | -10.2 | 0.6 |
| TC6 | Top of SIT Telescope Behind | -9.7 | BSIT.433 | -10.6 | 0.9 |
| TC7 | Top Front of SIT Telescope Behind | -11.0 | BSIT.425 | -11.8 | 0.8 |
| TC8 | Top of SIT Collimator Behind | -33.9 | BSIT.437 | -32.3 | 1.7 |
| TC9 | Right Side of SIT Collimator Behind | -32.0 | BSIT.419 | -31.9 | 0.1 |
| TC10 | Front of SIT Telescope Behind | -10.2 | BSIT.429 | -10.8 | 0.6 |
| TC11 | Top of SIT Telescope Behind | -9.1 | BSIT.433 | -10.6 | 1.5 |
| TC12 | Back of SIT Telescope Behind | -8.5 | BSIT.415 | -10.0 | 1.5 |
| TC13 | Left Side of SIT Ebox Behind | 0.9 | BSIT.431 | -1.6 | 2.5 |
| TC14 | Left Side of HVPS Behind | 0.5 | BPOWER.327 | -1.7 | 2.2 |
| TC15 | Left Side of SIT Telescope Behind | -8.4 | BSIT.413 | -10.2 | 1.9 |
| TC16 | Front of SIT Telescope Behind | -9.2 | BSIT.429 | -10.8 | 1.6 |
| TC17 | Back of SIT Telescope Behind | -8.5 | BSIT.415 | -10.0 | 1.5 |
| TC18 | Back of HVPS Behind | 0.1 | BPOWER.326 | -1.7 | 1.7 |
| TC19 | Back of SIT EBOX Behind | 0.4 | BSIT.409 | -1.7 | 2.1 |
| TC20 | Bottom of SIT EBOX Behind | 0.7 | BSIT.432 | -1.8 | 2.5 |
| TC21 | Bottom of SIT EBOX Behind | 0.3 | BSIT.432 | -1.8 | 2.1 |
| TC22 | Under SIT Collimator exterior Behind | -34.0 | BSIT.427 | -32.2 | 1.9 |
| TC23 | Right Side of SIT Collimator Behind | -32.9 | BSIT.419 | -31.9 | 1.0 |
| SEPT AHEAD NS |
| TC # | Location | Measured | NODE | Predicts | DELTA T |
| TC24 | Front of SEPTA on Bracket NS | -31.4 | SEPTNS.93 | -30.8 | 0.6 |
| TC25 | Front of SEPTA Telescope NS | -31.6 | SEPTNS.64 | -31.0 | 0.6 |
| TC26 | Front of SEPTA Telescope NS | -31.8 | SEPTNS.21 | -31.7 | 0.0 |
| TC27 | Front of SEPTA Telescope NS | -31.8 | SEPTNS.21 | -31.7 | 0.0 |
| TC28 | Front of SEPTA on Bracket NS | -30.8 | SEPTNS.106 | -30.9 | 0.1 |
| TC29 | Front of SEPTA on Bracket NS | -30.9 | SEPTNS.94 | -29.5 | 1.4 |
| TC30 | Front of SEPTA EBOX NS | -28.6 | SEPTNS.35 | -25.9 | 2.7 |
| TC31 | Back of SEPTA on Bracket NS | -31.1 | SEPTNS.105 | -30.8 | 0.3 |
| TC32 | Back of SEPTA Telescope NS | -31.3 | SEPTNS.82 | -31.0 | 0.3 |
| TC33 | Back of SEPTA Telescope NS | -32.1 | SEPTNS.28 | -31.7 | 0.4 |
| TC34 | Back of SEPTA Telescope NS | -32.1 | SEPTNS.28 | -31.7 | 0.4 |
| TC35 | Back of SEPTA on Bracket NS | -31.3 | SEPTNS.92 | -30.9 | 0.5 |
| TC36 | Back of SEPTA on Bracket NS | -30.2 | SEPTNS.109 | -29.5 | 0.7 |
| TC37 | Back of SEPTA EBOX NS | -28.1 | SEPTNS.40 | -25.9 | 2.2 |
| TC38 | Bottom of SEPTA Bracket NS | -28.8 | SEPTNS.109 | -29.5 | 0.7 |
| TC39 | Bottom of SEPTA Bracket NS | -29.7 | SEPTNS.94 | -29.5 | 0.2 |
| TC40 | Radiator Side of SEPTA NS | -28.8 | SEPTNS.39 | -26.1 | 2.7 |
| TC106 | Bottom of SEPTA NS | -29.0 | SEPTNS.36 | -25.2 | 3.8 |
| TC107 | Bottom of SEPTA NS | -27.8 | SEPTNS.36 | -25.2 | 2.6 |
| SEPT BEHIND NS |
| TC # | Location | Measured | NODE | Predicts | DELTA T |
| TC108 | Front of SEPTB on Bracket NS | -29.6 | SEPTNSB.93 | -31.7 | 2.1 |
| TC109 | Front of SEPTB Telescope NS | -29.8 | SEPTNSB.64 | -31.8 | 2.0 |
| TC110 | Front of SEPTB Telescope NS | -30.1 | SEPTNSB.21 | -32.5 | 2.4 |
| TC111 | Front of SEPTB Telescope NS | -29.9 | SEPTNSB.21 | -32.5 | 2.6 |
| TC112 | Front of SEPTB on Bracket NS | -29.0 | SEPTNSB.106 | -31.7 | 2.7 |
| TC113 | Front of SEPTB on Bracket NS | -28.3 | SEPTNSB.94 | -30.4 | 2.1 |
| TC114 | Front of SEPTB EBOX NS | -26.8 | SEPTNSB.35 | -26.8 | 0.0 |
| TC115 | Back of SEPTB on Bracket NS | -29.3 | SEPTNSB.105 | -31.6 | 2.4 |
| TC116 | Back of SEPTB Telescope NS | -29.7 | SEPTNSB.82 | -31.9 | 2.1 |
| TC117 | Back of SEPTB Telescope NS | -30.6 | SEPTNSB.28 | -32.6 | 2.0 |
| TC118 | Back of SEPTB Telescope NS | -30.7 | SEPTNSB.28 | -32.6 | 1.9 |
| TC119 | Back of SEPTB on Bracket NS | -29.3 | SEPTNSB.92 | -31.8 | 2.5 |
| TC120 | Back of SEPTB on Bracket NS | -30.0 | SEPTNSB.109 | -30.3 | 0.4 |
| TC121 | Back of SEPTB EBOX NS | -27.4 | SEPTNSB.40 | -26.8 | 0.6 |
| TC122 | Bottom of SEPTB Bracket NS | -28.6 | SEPTNSB.109 | -30.3 | 1.7 |
| TC123 | Bottom of SEPTB Bracket NS | -28.5 | SEPTNSB.94 | -30.4 | 1.9 |
| TC126 | Radiator Side of SEPTB EBOX NS | -27.3 | SEPTNSB.39 | -27.0 | 0.3 |
| TC124 | Bottom of SEPTB NS | -26.5 | SEPTNSB.36 | -26.1 | 0.3 |
| TC125 | Bottom of SEPTB NS | NA | SEPTNSB.36 | -26.1 | NA |

TABLE 3: Comparison of SIT AHEAD Test vs. Analysis Results (°C)

| Survival Balance |
| --- |
| TC # | Location | Measured | NODE | Predicts | DELTA T |
| TC1 | Front of SIT EBOX Ahead | -29.1 | ASIT.444 | -29.9 | 0.8 |
| TC2 | Front of HVPS Ahead | -28.5 | APOWER.325 | -29.6 | 1.1 |
| TC3 | Right Side of SIT EBOX Ahead | -28.1 | ASIT.431 | -29.8 | 1.7 |
| TC4 | Right Side of HVPS Ahead | -27.8 | APOWER.328 | -29.6 | 1.8 |
| TC5 | Right Side of SIT Telescope Ahead | -15.4 | ASIT.413 | -16.0 | 0.6 |
| TC6 | Top of SIT Telescope Ahead | -15.4 | ASIT.433 | -16.3 | 0.9 |
| TC7 | Top Front of SIT Telescope Ahead | -14.4 | ASIT.425 | -16.0 | 1.6 |
| TC8 | Top of SIT Collimator Ahead | -12.7 | ASIT.437 | -14.5 | 1.8 |
| TC9 | Lengthened Side SIT Collimator Ahead | -15.5 | ASIT.417 | -13.8 | 1.7 |
| TC10 | Front of SIT Telescope Ahead | -17.3 | ASIT.429 | -15.9 | 1.4 |
| TC11 | Top of SIT Telescope Ahead | -15.8 | ASIT.433 | -16.3 | 0.5 |
| TC12 | Back of SIT Telescope Ahead | -16.7 | ASIT.415 | -16.2 | 0.5 |
| TC13 | Left Side of SIT EBOX Ahead | -29.9 | ASIT.445 | -30.2 | 0.3 |
| TC14 | Left Side of HVPS Ahead | -28.7 | APOWER.327 | -29.6 | 0.9 |
| TC15 | Left Side of SIT Telescope Ahead | -18.5 | ASIT.435 | -16.5 | 2.0 |
| TC16 | Front of SIT Telescope Ahead | -17.8 | ASIT.429 | -15.9 | 1.9 |
| TC17 | Back of SIT Telescope Ahead | -16.9 | ASIT.415 | -16.2 | 0.7 |
| TC18 | Back of HVPS Ahead | -28.0 | APOWER.326 | -29.6 | 1.7 |
| TC19 | Back of SIT EBOX Ahead | -29.1 | ASIT.409 | -29.9 | 0.8 |
| TC20 | Bottom of SIT EBOX Ahead | -29.6 | ASIT.432 | -29.8 | 0.2 |
| TC21 | Bottom of SIT EBOX Ahead | -32.9 | ASIT.432 | -29.8 | 3.1 |
| TC22 | Under SIT Collimator exterior Ahead | -12.6 | ASIT.427 | -12.1 | 0.5 |
| TC23 | Heater Side SIT Collimator Ahead | -4.0 | ASIT.419 | -5.3 | 1.3 |
| Hot Operational Balance |
| TC # | Location | Measured | NODE | Predicts | DELTA T |
| TC1 | Front of SIT EBOX Ahead | 3.2 | ASIT.444 | 5.1 | 1.9 |
| TC2 | Front of HVPS Ahead | 3.1 | APOWER.325 | 5.1 | 2.0 |
| TC3 | Right Side of SIT EBOX Ahead | 4.7 | ASIT.431 | 5.3 | 0.6 |
| TC4 | Right Side of HVPS Ahead | 3.6 | APOWER.328 | 5.2 | 1.6 |
| TC5 | Right Side of SIT Telescope Ahead | -4.1 | ASIT.413 | -0.6 | 3.5 |
| TC6 | Top of SIT Telescope Ahead | -3.4 | ASIT.433 | -0.9 | 2.5 |
| TC7 | Top Front of SIT Telescope Ahead | 0.4 | ASIT.425 | -0.4 | 0.7 |
| TC8 | Top of SIT Collimator Ahead | 6.6 | ASIT.437 | 3.9 | 2.7 |
| TC9 | lengthed Side SIT Collimator Ahead | 4.8 | ASIT.417 | 4.6 | 0.2 |
| TC10 | Front of SIT Telescope Ahead | -2.9 | ASIT.429 | -0.4 | 2.5 |
| TC11 | Top of SIT Telescope Ahead | -4.6 | ASIT.433 | -0.9 | 3.7 |
| TC12 | Back of SIT Telescope Ahead | -4.5 | ASIT.415 | -0.9 | 3.6 |
| TC13 | Left Side of SIT EBOX Ahead | 1.8 | ASIT.445 | 4.7 | 2.9 |
| TC14 | Left Side of HVPS Ahead | 2.6 | APOWER.327 | 5.1 | 2.5 |
| TC15 | Left Side of SIT Telescope Ahead | -4.4 | ASIT.435 | -1.3 | 3.1 |
| TC16 | Front of SIT Telescope Ahead | -3.7 | ASIT.429 | -0.4 | 3.3 |
| TC17 | Back of SIT Telescope Ahead | -4.2 | ASIT.415 | -0.9 | 3.3 |
| TC18 | Back of HVPS Ahead | 3.2 | APOWER.326 | 5.1 | 2.0 |
| TC19 | Back of SIT EBOX Ahead | 3.6 | ASIT.409 | 5.1 | 1.4 |
| TC20 | Bottom of SIT EBOX Ahead | 2.8 | ASIT.432 | 5.3 | 2.5 |
| TC21 | Bottom of SIT EBOX Ahead | -0.4 | ASIT.432 | 5.3 | 5.7 |
| TC22 | Under SIT Collimator exterior Ahead | 6.8 | ASIT.427 | 6.5 | 0.3 |
| TC23 | Heater Side SIT Collimator Ahead | 17.3 | ASIT.419 | 15.3 | 1.9 |

Thermal Vacuum Results

The TV test consisted of six operational qualification cycles, a cold turn-on, a hot turn-on and both a hot/cold survival soak were accomplished. The SIT Behind telescope surface barrier detector was not working correctly and a replacement was not available. However, the TV test was carried out testing the function of all but the failed detector on the SIT Behind instrument. Additional testing of SIT Behind will occur when a replacement detector becomes available. The SIT Ahead instrument has satisfactorily completed the thermal vacuum test program and has met all of the requirements spelled out in the STEREO-GSFC-002.C test plan and STEREO Environment Definition, Observatory, Component and Instrument Test Requirements Document.

The test setup was carried out according to the STEREO-GSFC-002.C test plan. Two units were mounted side by side on a heat exchanger (see Figure 6).

To provide good thermal contact to the cold plate the flight Ultem stand-offs below the mounting feet were omitted from the bottom plate of the E-Boxes.

Figure 6 SIT TV Test Set-up

1 = Stabilize, Soak 1Hour and Test Concurrently.

2 = Stabilize, Turn On, Soak 1Hour and Test Concurrently.

Figure 7: SIT Thermal Vacuum Test Temperature Profile

Figure 8: SIT Thermal Vacuum Test Temperature Data

Table 5: SIT Thermal Vacuum Soak Temperatures

| Component | Bakeout (OC) | Hot Srv (OC) | Hot Op (OC) | Cold Srv (OC) | Cold Op (OC) |
| --- | --- | --- | --- | --- | --- |
| SIT EBOX | 50 | 50 | 45 | -40 | -30 |
| SIT Telescope | 32 | 32 | 30 | -40 | -30 |

Conclusions

- The SIT Ahead instrument has satisfactorily completed the thermal vacuum test program.

- The SIT Behind instrument will undergo addition TV testing, to complete the thermal vacuum test program, when the surface barrier detector is replaced.

- The thermal balance test correlation has successfully verified the SIT instruments thermal models and the thermal conductance of the isolators.

- Thermal coatings chosen for the SEPT/SIT instruments are adequate for maintaining temperatures in the acceptable range.

- The SIT qualification test has been conducted from –40ºC to +50ºC (survival), which indicates significantly more margin at the hot end than for the cold case for the SIT Ahead instrument. When making the SIT Ahead instrument’s flight MLI blanket, the radiator area was reduced to provide additional heater margin at the cold end with the minimum expected flight voltage.

- The SIT qualification test has been conducted from –30ºC to +30ºC (operational), which for the SIT Ahead instrument, again indicates significantly more margin at the hot end than for the cold case. The reduced radiator area will provide additional heater margin at the cold end.

- All the SEPT units proved to have adequate margin at both the hot and cold ends.

# Text boxes

Calrod Circuit 7

Calrod Circuit 6

B_SEPTE

A_SEPTE

A_SIT

Mounting

Plate W/MLI

Shield W/MLI

on back

A_SEPTNS

B_SEPTNS

B_SIT

Mounting Plate W/MLI